# Lung Segmentation Session 9 Report

Preliminary internal evaluation of a U-Net model for lung-field segmentation in chest X-rays.

# Project status

This phase packages the current training checkpoint, evaluation artifacts, and code for a private research repository. It does not include a revised data split or retraining. The work is an educational/research prototype and is not clinically validated.

# Dataset quality and empty masks

The Darwin subset contains 6,106 image-mask pairs. Inspection identified 81 empty masks using the current foreground rule (mask pixel value >127); these samples are absent from the recorded split. The fixed split contains 6,025 unique pairs: 4,820 training, 602 validation, and 603 test. A direct check found no overlap among partitions, no split filenames missing from the paired dataset, and none of the 81 empty-mask samples in any partition. The reasons for empty masks have not been fully investigated; their exclusion criterion and implications remain part of the data-quality audit.

# Evaluation results

The selected checkpoint was evaluated on the fixed held-out test list using a 0.5 probability threshold. Results are preliminary and internal to the current dataset and workflow.

| Metric | Mean | Median | SD | N |
| --- | --- | --- | --- | --- |
| Dice | 0.9470 (94.70%) | 0.9562 | 0.0318 | 603 |
| IoU | 0.9010 (90.10%) | 0.9161 | 0.0543 | 603 |

The summary reports zero empty ground-truth masks. Per-image results and the worst ten cases are included in Results/. The worst cases show substantial false-positive and false-negative regions. No confidence intervals have been calculated.

# Method summary

The pipeline uses grayscale 512×512 images normalized to [0,1], binary masks resized with nearest-neighbor interpolation, and a 2D encoder-decoder U-Net with skip connections and sigmoid output. Training scripts use binary cross-entropy plus soft Dice loss, Adam at an initial learning rate of 1e-4, and batch size 2 for the Session 9 runs. The fixed split contains 4,820 training, 602 validation, and 603 test images.

# Interpretation and limitations

This is an internal evaluation; independent external-dataset testing has not been documented.

Patient-level grouping, duplicate detection, and leakage across partitions have not yet been audited.

A random image-level partition may put related images from one patient or source in multiple subsets.

The reported metrics do not demonstrate clinical utility, diagnostic accuracy, or safe use in care.

Source dataset provenance, label provenance, and redistribution permissions need to be documented before any public release.

# Next research steps

Audit dataset sources, licenses, duplicates, patient/study identifiers, and split leakage.

Freeze the evaluation protocol and document exact preprocessing, metric aggregation, threshold selection, checkpoint selection, and software versions.

Compute confidence intervals and subgroup/error analyses using appropriate independent units.

Test the frozen model on an independently sourced, appropriately licensed dataset.

Compare against suitable baselines, evaluate robustness, conduct a literature review, and define a specific defensible research question before making paper claims.

# Reproducibility notes

The current scripts contain machine-specific absolute paths and expect the source images and masks in Darwin/img and Darwin/mask. Dataset files are intentionally excluded from this repository package. The split manifest contains image filenames; review it for privacy and redistribution concerns before adding it to a public repository. The model weights are included for this private repository draft; verify GitHub file-size handling and access controls before upload.

Use restriction: research and education only. Not for diagnosis, treatment, triage, or clinical decisions.